Anesthesia Machine

| No. | Technical Specifications | Min KAUH Requirement |
| --- | --- | --- |
| 1 | Manufacturer | Please specify manufacturer and country of origin |
| 2 | Model Number | Please specify model number of the offered equipment |
| 3 | Safety standard | FDA approval or CE marking |
| 4 | Anesthesia Gas Delivery System | Should have provision for delivery of oxygen, nitrous oxide and medical air with pressure gauges. Should have independent attachments for connecting central gas supply and pin indexed cylinders. Oxygen and nitrous oxide should be linked either mechanically or pneumatically to ensure a minimum of 25% oxygen delivery at all times to avoid delivery of hypoxic mixture. Should have audio-visual oxygen failure warning system with nitrous oxide cut off. Should have back bar which is ISO pin type to attach vaporizer easily. Should be supplied with necessary attachments for use of the breathing circuits (bains, Jackson-rees, and closed circuit etc…). Should have top shelf to keep monitors and a tabletop to keep anesthetic drugs, equipments etc…. The machine should possess battery back up for electrical components. Castor wheels should be durable and moisture resistant and smooth. The anesthesia machine frame should be made of rust proof material/stainless steel. |
| 5 | Standard Circle Absorber System | Should have adjustable pressure limiting valve, breathing circuit pressure measuring device. Should have a bag/ventilator selecting valve integrated onto the absorber. Should be suitable to use low flow techniques. Facility to attach oxygen sensor. Should have CO2 absorbent chamber canister. |
| 6 | Precision Vaporizers (Temperature, Pressure, and Flow compensated) for Halothane, Isoflurane and Sevoflurane | Should be easy to amount and dismount from the back bar. Vaporizers should have ISO pin type (select atec) mounting and vaporizer interlocking facility. Should have a standard filling port with keyed filling device. Should designed for transport with liquid in vaporizer chamber with protection against tipping and shaking. Maintenance free vaporizer. Should be a bag in bottle anesthesia ventilator with standing (ascending/Piston) bellows. |
| 7 | Ventilator | Should be supplied with Adult and Pediatric bellows. Should be able to set tidal volume, respiratory rate and I:E ratio. Ventilator should have audible alarms for ventilator failure, low oxygen supply pressure, inadequate volume delivery, disconnecting alarm, and power supply failure. Should have battery backup for 30min at minimum. |
| 8 | Monitoring System should provide facility to monitor | Oxygen and nitrous oxide and anesthetic agent in the inspired and expired mixture. Inspired and end tidal carbon di-oxide through side stream. Airway pressure monitoring should be present. Glare free colored monitors with large screen for easy visibility. Ventilator screen: Display screen size: at least 12 inch with high resolution, full color. |
| 9 | General | Integrated suction unit with 2 jars: vacuum and or oxygen. Drawer: at least one. Demonstration is must. |
| 10 | Main Power | 220 - 240 VAC , 50 Hz |

Patient Monitor

| No. | Technical Specifications | Min KAUH Requirement |
| --- | --- | --- |
| 1 | Manufacturer | Please specify manufacturer and country of origin |
| 2 | Model Number | Please specify model number of the offered equipment |
| 3 | Safety standard | FDA approval or CE marking |
| 4 | Monitoring application | Neonatal, Pediatric, and Adult |
| 5 | Configuration | Wall Mount, Modular,icu monitor with intensive care software. Modular type. Multi parameter vital sign module. Module are freely exchangeable between all monitors. |
| 6 | Display Parameters: | Parameters: ECG, Respiration rate,NIBP,Body tempreture, spo2 ( technology to be mentioned),IBP. *Pulse Oximetry SPO2 (1-100%). *Non- Invasive Blood Pressure (NIBP), Adult Cuff (Medium and Large Size) and Two Sizes of Pediatric Cuffs. *Invasive Blood Pressure (IBP). *Temperature Measurement. *EtCO2 Module with all Accessories (Optional). *Cardiac Output.(optional) *EEG Module with all Accessories (Optional). *BIS/Entropy Module (Optional). At least 8 Digital and waves/ traces Display. *Audio Alarm System. *High resolution TFT-LCD 15”( minimum),touch screen *Necessary Mounting Solution/Mounting on any Pendant for Monitors. *Transport Monitor with at least 30 minutes battery back up. |
| 7 | Main Power | 220 - 240 VAC , 50 Hz |

Anesthesia Trolley

- Trolley manufactured from aluminium profiles, closed on three side with laminate .Top plate in stainless steel.

- Support Rails on both sides.

- Five transparent boxes at upper part of trolly with two needle disposal holder.

- At least fuor easy running drawers included deviders, with lock

- Waste bin

- IV stand

- Frame with tp-out bins

- 4 swivel castors with brake