Anesthesia Machine

| No. | Technical Specifications | Min KAUH Requirement |
| --- | --- | --- |
| 1 | Manufacturer | Please specify manufacturer and country of origin |
| 2 | Model Number | Please specify model number of the offered equipment |
| 3 | Safety standard | FDA approval or CE marking |
| 4 | Anesthesia Gas Delivery System | - Should have provision for delivery of oxygen, nitrous oxide and medical air with pressure gauges. - Should have independent attachments for connecting central gas supply and pin indexed cylinders. - Oxygen and nitrous oxide should be linked either mechanically or pneumatically to ensure a minimum of 25% oxygen delivery at all times to avoid delivery of hypoxic mixture. - Should have audio-visual oxygen failure warning system with nitrous oxide cut off. - Should have back bar which is ISO pin type to attach vaporizer easily. - Should be supplied with necessary attachments for use of the breathing circuits (bains, Jackson-rees, and closed circuit etc…). - Should have top shelf to keep monitors and a tabletop to keep anesthetic drugs, equipments etc…. - The machine should possess battery back up for electrical components. - Castor wheels should be durable and moisture resistant and smooth. - The anesthesia machine frame should be made of rust proof material/stainless steel. |
| 5 | Standard Circle Absorber System | - Should have adjustable pressure limiting valve, breathing circuit pressure measuring device. - Should have a bag/ventilator selecting valve integrated onto the absorber. - Should be suitable to use low flow techniques. - Facility to attach oxygen sensor. - Should have CO2 absorbent chamber canister. |
| 6 | Precision Vaporizers (Temperature, Pressure, and Flow compensated) for Halothane, Isoflurane and Sevoflurane | - Should be easy to amount and dismount from the back bar. - Vaporizers should have ISO pin type (select atec) mounting and vaporizer interlocking facility. - Should have a standard filling port with keyed filling device. - Should designed for transport with liquid in vaporizer chamber with protection against tipping and shaking. - Maintenance free vaporizer. - Should be a bag in bottle anesthesia ventilator with standing (ascending/Piston) bellows. |
| 7 | Ventilator | - Should be supplied with Adult and Pediatric bellows. - Should be able to set tidal volume, respiratory rate and I:E ratio. - Ventilator should have audible alarms for ventilator failure, low oxygen supply pressure, inadequate volume delivery, disconnecting alarm, and power supply failure. - Should have battery backup for 30min at minimum. |
| 8 | Monitoring System should provide facility to monitor | - Oxygen and nitrous oxide and anesthetic agent in the inspired and expired mixture. - Inspired and end tidal carbon di-oxide through side stream. - Airway pressure monitoring should be present. - Glare free colored monitors with large screen for easy visibility. - Ventilator screen: - Display screen size: at least 12 inch with high resolution, full color. |
| 9 | General | - Integrated suction unit with 2 jars: vacuum and or oxygen. - Drawer: at least one. - Demonstration is must. |
| 10 | Main Power | 220 - 240 VAC , 50 Hz |

Patient Monitor

| No. | Technical Specifications | Min KAUH Requirement |
| --- | --- | --- |
| 1 | Manufacturer | Please specify manufacturer and country of origin |
| 2 | Model Number | Please specify model number of the offered equipment |
| 3 | Safety standard | FDA approval or CE marking |
| 4 | Monitoring application | Neonatal, Pediatric, and Adult |
| 5 | Configuration | - Wall Mount, Modular,icu monitor with intensive care software. - Modular type. - Multi parameter vital sign module. - Module are freely exchangeable between all monitors. |
| 6 | Display Parameters: | Parameters: ECG, Respiration rate,NIBP,Body tempreture, spo2 ( technology to be mentioned),IBP. *Pulse Oximetry SPO2 (1-100%). *Non- Invasive Blood Pressure (NIBP), Adult Cuff (Medium and Large Size) and Two Sizes of Pediatric Cuffs. *Invasive Blood Pressure (IBP). *Temperature Measurement. *EtCO2 Module with all Accessories (Optional). *Cardiac Output.(optional) *EEG Module with all Accessories (Optional). *BIS/Entropy Module (Optional). At least 8 Digital and waves/ traces Display. *Audio Alarm System. *High resolution TFT-LCD 15”( minimum),touch screen *Necessary Mounting Solution/Mounting on any Pendant for Monitors. *Transport Monitor with at least 30 minutes battery back up. |
| 7 | Main Power | 220 - 240 VAC , 50 Hz |

Anesthesia Trolley

- Trolley manufactured from aluminium profiles, closed on three side with laminate .Top plate in stainless steel.

- Support Rails on both sides.

- Five transparent boxes at upper part of trolly with two needle disposal holder.

- At least fuor easy running drawers included deviders, with lock

- Waste bin

- IV stand

- Frame with tp-out bins

- 4 swivel castors with brake