المواصفات الفنية قسم غسيل الكلى

Technical Specifications for HemoDialysis Unit

| No. | Technical Specifications | #### Min KAUH Requirement |
| --- | --- | --- |
| 1. | Manufacturer | Please specify manufacturer and country of origin & country of manufacturer |
| 2. | Model number | Please specify model number |
| 3. | Safety standard | FDA Approval or CE marking |
| 4. | Dialysate delivery : - Degassing | Included |
|  | - Proportioning system | Volumetric or servo-controlled |
|  | - Temperature control (C°) | 35 to 39 |
|  | - Temperature alarm limit (C°) | ± 1 C° of set temperature |
|  | - Conductivity range (mS/cm) | 12.5 to 15.5 approx. |
|  | - Flow (ml/min) | 300 to 700 |
|  | - Trans membrane pressure (mmHg) | Approx (0 to +450) |
|  | - Sodium therapy (Profile) | Included |
|  | - Bicarbonate dialysis | Included |
|  | - Ultra filtration removal (L/hr) | 0 to 3, include (UF profile) |
| 5. | Bypass indicator | Included |
| 6. | Blood leak detector | Included |
| 7. | Blood circuit : - Venous clamp | Included |
|  | - Arterial pressure reading (mmHg) | Approx (-300 to +150) |
|  | - Venous pressure reading (mmHg) | Approx (0 to +400) |
|  | - Blood pump range (ml/min) | 50 to 500 |
|  | - Heparin pump range (ml/hr) | 0.1 to 10 approx. |
|  | - Air / foam detector | Included |
| 8. | Disinfection program | - Rinse - Hot rinse is preferable - Disinfection and decalcification with chemicals It must be mandatory rinse after each disinfection process , even , in case of power failure during process |
| 9. | Displayed parameters : - Display Type - Dialysate pressure - Temperature - Transmembrane pressure - Conductivity - Flow rate - Elapsed time - Remaining time - Other parameters - Language | LCD or CRT is preferable Preferable Included Included Included Included )Preferable) Included Please specify if there are more Specify |
| 11. | Internal Diameter of blood pump segment | Please specify the range |
| 11. | Internal Diameter of blood pump segment | Please specify the range |
| 12. | Alarm actions : - Blood leak detector - Air / Foam detector - Arterial / Venous pressure (high / low) - Temperature (high / low) - Conductivity (high/ low) | - Stop blood pump , clamp line & alarm - Stop blood pump , clamp line & alarm - Stop blood pump & alarm - Bypass and alarm - Bypass and alarm |
| 13. | Operation | - Compulsory self test (Diagnostic program) before each dialysis |
| 14. | Battery Backup | - Included - Auto Battery Operation |
| 15. | Application | - Acetate and bicarbonate mode - Double needles |
| 16. | System | It should be open system , operate with all types dialyzers and disposables , in case of closed system refer to special conditions |
| 17. | Other option | Mention and specify price |
| 18. | Blood pressure monitor | Included |
| 19. | On Line clearance monitor (KT/V) | Included/built in device for measurement monitoring effective urea clearance and dialysis dose (K/V) |
| 20. | Construction | - Mobile - 4 Antistatic and swiveled castors |
| 21. | Maintenance software | Allow connecting PC to machine by interface socket and make diagnostic for all parts ( troubleshooting ) |
| 22. | Consumable | All consumables required for installation and standardization of system (included bacterial and water filter to be replaced according to manufacturer recommendation) and to be given for free of cost as needed. |
| 23. | Power supply | 220 V ± 5% / 50 Hz |

Un code : 42161632

الموضوع: المواصفات/ كراسي لوحدة غسيل الكلى

Technical Specifications for Hemodialysis Chair

| No. | Technical Specifications | - KAUH Requirement |
| --- | --- | --- |
|  | Manufacturer | Please specify manufacturer and country of origin & country of manufacturer |
|  | Model number | Please specify model number |
|  | Safety standard | FDA Approval or CE marking |
|  | Construction | -the chair must not contain any wooden parts -leather or water proof and cover all mattress - Frame material of steel or aluminum covered with bacteriological, chemical and fungal resistant material. - easy to clean and sterilize. -motors and electrical parts should be isolated and covered properly. |
|  | CPR | INCLUDED |
|  | Remote Control | INCLUDED |
|  | Backrest (with head pillow), foot rest, and height should be adjusted independently | Electrically (by remote) |
|  | Trendelenburg position | Included |
|  | Arm rest | Horizontally and vertically movement |
|  | Enough length for whole arm | Not less than (50) cm |
|  | Dimension: enough length, height and width for all weight | Not less than (210*50*55) |
|  | Weight load | Not less than 200 Kg |
|  | castors | 4 antistatic castors and all swivel with two lockable (break) at least |
|  | IV pole mount, weighting scale, dampers | Preferred |
|  | Rechargeable Battery | Preferred |
|  | Power supply | 220 V± 5%/50 Hz |

الموضوع: المواصفات الفنيه / كراسي وحدة العلاج الكيماوي

Technical Specifications for Chemotherapy Chair

| No. | Technical Specifications | - KAUH Requirement |
| --- | --- | --- |
| 1. | Manufacturer | Please specify manufacturer and country of origin & country of manufacturer |
| 2. | Model number | Please specify model number |
| 3. | Safety standard | FDA Approval or CE marking |
| 4. | Construction | -the chair must not contain any wooden parts -leather or water proof and cover all mattress - Frame material of steel or aluminum covered with bacteriological, chemical and fungal resistant material. - easy to clean and sterilize. -motors and electrical parts should be isolated and covered properly. |
| 5. | CPR | INCLUDED |
| 6. | Remote Control | INCLUDED |
| 7. | Backrest (with head pillow), foot rest, and height should be adjusted independently | Electrically (by remote) |
| 8. | Trendelenburg position | Included |
| 9. | Arm rest | Horizontally and vertically movement |
| 10. | Enough length for whole arm | Not less than (50) cm |
| 11. | Dimension: enough length, height and width for all weight | Not less than (210*50*55) |
| 12. | Weight load | Not less than 200 Kg |
| 13. | castors | 4 antistatic castors and all swivel with two lockable (break) at least |
| 14. | IV pole mount, weighting scale, dampers | Preferred |
| 15. | Rechargeable Battery | Preferred |
| 16. | Power supply | 220 V± 5%/50 Hz |

Un code : 42161605