Technical Specifications for Node Seeker (Gamma Probe)

| No. | Technical Specifications | Min KAUH Requirement |
| --- | --- | --- |
| 1 | Manufacturer | Please specify manufacturer and country of origin |
| 2 | Model Number | Please specify model number of the offered equipment |
| 3 | Safety Standard | FDA approval or CE marking |
| 4 | Main Features | Node Seeker: - Radiation detection Probe for identification of radiolabeled tissue for localization of sentinel lymph nodes in patients with breast cancer and melanoma. -detection of gamma or beta rays. -hand held use. - a collimator allows entry of gamma rays only from the front of the probe. Gamma Probe Specification: - Angular resolution : 42 degree or less. - Spatial resolution: 13mm or less. - Lateral shielding, Tc-99m: at less 99%. - Absolute to (Tc-99m) : equal or less than 30000 cps/MBq. - Energy Thresholds: Adjustable above 110 KeV for 99mTC and other radionuclide . Control Unit Specifications: -Count Display: CPS refreshed as low as every 0.1 sec . -Energy Discrimination: available. -Isototope Selection: Preset with variable windows. -Audio Guidance: available. System Sterilizable or Sterile disposable covers are provided by vendor. |
| 5 | Quality Assurance | Automated QA Program with QC data retention |
| 6 | Main Power | 220 - 240 VAC , 50 Hz |