| No. | Technical Specification | Min Requirement | Bidder Offer |
| --- | --- | --- | --- |
| 1 | Manufacturer | Please specify Manufacturer and country of origin (USA, Europ, Japan) |  |
| 2 | Model Number | Please specify model number of the offered equipment |  |
| 3 | Safety Standard | FDA or CE marking |  |
| 4 | Desine and Quality | Compact heavy duty and high quality |  |
| 5 | Ventilator System | - Microprocessor controlled used for Adult, Pediatric - Software must be upgradable - Humidifier - Compressor (optional) - Nebulizer (standard) - Leak compensation |  |
| 6 | Ventilator mode | - Volume controlled - Pressure controlled (APRV) - SIMV - Spontaneous pressure support/ CPAP - Non invasive - PRVC |  |
| 7 | Special functions | - Pediatric , Adult applecation - Apnea back up ventilation - Standby - Manual breaths - Measure static and dynamic parameter - High frequency ventilation mode (optional) |  |
| 8 | Graphic waveform, trends & loops Minimume 3 waveform | - Flow - Volume - Pressure - Other specify |  |
| 9 | Display | Multicolor LCD display not less than "10" Preferable touch screen With easy to use inter face |  |
| 10 | Ventilator monitoring parameters | - Inspiratory peak flow - Expired tidal volume - Pressure (peak inspiratory,mean, peep) - Inspired oxygen concentration Fio2 - I:E ratio - frequency |  |
| 11 | Trigger mechanism | - pressure, flow |  |
| 12 | Alarm limits (adjustable) | - low/high pressure limit - low Fio2 - low/high PEEP/CPAP pressure - High /low tidal volume - low minute volume - low/high respiratory rate - low/high Fio2 :21-100% - Apnea: 5-20 sec. - All alarms audible and visible |  |
| 13 | - Power supply | - 220v 50Hz - (built in)Battery back up not less than 30min |  |
| 14 | Equipment alarm | - Power failure - Low Battery - Disconnection - Gas supply failure - Flow sensore alarms - All alarms audible and visible |  |
| 15 | - Self diagnostic error/calibration massage | - Included |  |