| No. | Technical Specifications | Min KAUH Requirement | Compley | Notes |
| --- | --- | --- | --- | --- |
| 1 | Manufacturer | Please specify manufacturer and country of origin |  |  |
| 2 | Model Number | Please specify model number of the offered equipment |  |  |
| 3 | Safety Standard | FDA approval or CE marking |  |  |
| 4 | Main features | - Fully automated ,microprocessor controlled with access operator control panel ,touch screen operation - Should perform single or double access aphaeresis , in single needle procedure the equipment should continue to process the whole blood during the return cycle continue flow procedure in single access. - Capable of doing prime with normal saline and or mixture or normal saline and ACD. Integrated blood pressure monitor Integrated blood warmer Hemlysis detector Collection should be fully automatic Temperature sensor Automatic valves Color coded blood pumps Control panel with anticoagulation pump,acd drip monitor and air detector. Lower extra corporeal volume, more than 200 to 250 ml in case of both single and double needle aphaeresis Advance plasma weigher Capable of down loading or printing full procedure repory any time after procedure. Continuous flow device Inlet and outlet pressure monitor Automated safety check prior to procedure mandatory Fail safe electronic circuit for top priority safety system Return line air detector Replacement fluid detector Supervision of extracorporeal volume Return line clamp with optical detector for blood detection. Bar code reader :optional Continuous monitoring of collection to avoid any contamination through interface detecto |  |  |
| 5 | technology | Centrifuge or membrane - Centrifugal separation - Centrifuge compartment - Short dwell time inside the centrifuge - Components collected outside centrifuge - Leakage detector inside the centrifuge. |  |  |
| 6 | The following options should be included | All consumable required for installation and standardization of system should be given free of charge. |  |  |
| 7 | Main Power | 220 - 240 VAC, 50 Hz Suitable UPS with minimum one hour back up battery |  |  |