## العرض الفني
| Specefcation | VENDER ITEM DESCRIPTION | COMMENTS | WHERE TO FIND
|  |  |  | OFFER PAGE NO | CATALOUGE PAGE NO
Portable RO | pre filter to remove particulates greater than 5 microns
Stainless steel pump and automatic flushing
pressure adjust regulator and gauges
inlet water sensor to monitors of the inlet water
Automatic softener and Automatic carbon filter compatible with RO system
disinfectant port
TDS display monitor and pressure gauges
RO pump power switches s mounted on the front panel
FDA registered manufacturer meats stringent medical devces requirements
water specification required | TDS: not more than 5ppm
TH: Zero
production rate >75 L/hr
Hospital water | TDS:400 ppm
TH:100 ppm
CI: 0.4 ppm
على الشركة تعبئة هذا الكشف وبيان مدى موافقة الاجهزة المعروضة من قبلة للمواصفات المطلوبة وبيان المعلومات التي تثبت ذلك بحيث يتم وضع علامة (X) أمام المواصفة اذا كانت مطابقة للمطلوب و(XX) اذا كانت غير مطابقة وحسب الأعمدة أعلاه.
على الشركة التقدم بعرض مالي منفصل للأجهزة المطلوبة.