- General Specifications

  - The navigation system shall support:

    - Trans-cranial tumor resection

    - Frameless stereotactic biopsy procedures

    - Trans-nasal endoscopic FESS (Functional Endoscopic Sinus Surgery) and Skull base surgeries

  - Shall provide full instruments set for:

    - Brain Tumor resection (i.e. Articulated arm, two patient frames, straight probe, registration probe, calibrator, sterilization tray)

    - Biopsy procedures (i.e. Articulated arm, dual adaptor, biopsy aiming device, biopsy probe, reducing tube, sterilization tray)

    - Trans-nasal procedures (i.e. Bayonet probe, registration probe, maxillary and frontal seekers, forehead patient reference, universal tracker, calibrator, sterilization tray).

  - The navigation cart shall be:

    - Stand-alone with only one power cable (No cable running to the patient table).

    - A mobile touch display cart and a separate infrared camera cart for flexible positioning in the O.R.)

    - Small Display Cart footprint, shall be equal or less than 55 x 45 cm

    - Motorized joints for remote-controlled camera alignment with integrated laser pointer for fast and intuitive camera alignment.

    - Camera Height: Range from 1000 mm to 2485 mm.

    - Easily accessible connection panel for integration of intraoperative devices.

    - 100-V – 240V, 50-60Hz.

  - Shall satisfy the minimum requirements for the navigation computer:

    - Large 32" capacitive touch display for clear viewing and easy control

4K resolution (3840 x 2160 pixels) for brilliant visualization of clinical data

    - Processor: Intel Core™ i9 10-core Processor 3.7/4.7 GHz, 32 GB RAM memory and 1 TB SSD

    - Peripherals: Minimum 2 x USB 3.0, Ethernet

    - Glass front for easy cleaning and disinfection.

  - The registration process shall take no longer than 2 minutes, and shall preferably include contactless acquisition of surface registration points on patients’ skin.

  - Shall provide options to mount patient reference frame on:

    - The skull holder

    - The patient forehead

  - Shall support CT & MRI images with image fusion capability, with ultrasound digital integration capability.

  - Import DICOM images from DVD, USB and PACS.

  - Shall support anatomical landmark registration and trace/surface registration.

  - Shall provide planning features for segmentation of tumor and 3D modeling.

  - Shall provide planning features for virtual craniotomy and visualizing brain tissues and fiber tracking capabilities.

  - Shall provide features for compensation for distortions in MR datasets with deformable co-registration enabling contouring and tractography across modalities.

  - Shall support virtual tip extension.

  - Shall provide Navigation layouts to display anatomical, trajectory and probe eye views.

  - Shall support automatic registration using any intra-op imaging modality.

  - Shall have video documentation of the entire navigation procedure.

# Tracking Device

  - Tracking system shall use the optical technology with accuracy of 0.2mm or better.

  - The Tracking system shall provide reliable accuracy and will not be interfered in presence of any metallic surgical instrument or under OR light.

  - Instruments reflective markers designed to be disposable.

  - Instruments will be pre-calibrated and will not require re-calibration before each case.

  - Shall provide a minimum of Two (2) expandable universal trackers to transform any surgical instruments to navigation enabled instruments.

  - Shall provide visual feedback of the accuracy of the instruments tips and shall calibrate them if they are out of calibration.

  - Shall provide live video from the working field for intuitive positioning of the tracking device and avoiding line of sight issues.

# Consumables

  - Shall provide free supply of 15 disposable biopsy needles.

# Regulatory, Warranty and Support

  - Shall be CE ap proved.

  - Shall provide 2 years’ comprehensive warranty from the date of installation and applied use.

  - Shall provide on-site training of the system to the OR staff.

  - Shall provide free software updates during the warranty period

  - Shall provide free online and remote support during the product’s lifetime.

Optional:

- The navigation system supports spine procedures.

- Provide full instruments set for Spine Procedures (i.e. navigated awl, registration probe, patient reference frame, spinous clamp, Percutaneous Surgery Unit, Universal trackers, calibrator, automatic registration solution for intra-op 3D CT, sterilization tray) C-Arm Registration.

- Shall provide options to mount patient reference frame on attachment for Open or MIS Spine surgery

- Shall provide free supply of disposable parts for spine surgeries for 20 procedures.

- Shall support pre-op and intra-op pedicle screw planning.