Technical Specifications for Lithotripsy Pneumatic and Ultrasonic Machine

| No. | Technical Specifications | Min KAUH Requirement |
| --- | --- | --- |
| 1 | Manufacturer | Please specify manufacturer and country of origin |
| 2 | Model Number | Please specify model number of the offered equipment |
| 3 | Safety Standard | FDA approval or CE marking |
| 4 | Technology | Ultrasound ,pneumatic and combination Dual energy in one handpiece and one probe |
| 5 | Main Features | -main application is stone fragmentation in Kidney, Ureter and bladder. -Friendly user interface. Three separated modes: ultrasonic, Pneumatic, and combination, combined energy in a single probe designe - ultrasonic , pneumatic , and suction options. - possibility of use of impact and ultrasounds independently and in combination. - ballistic impact generated electro magnetically with a coil - ultrasound with auto adaptive algorithm. -adjustable pneumatic energy up to (80mJ) approx. or better, specify. -adjustable pneumatic frequency (0-10) Hz approx. or better, specify. -adjustable ultrasonic power up to (70W) approx. or better, specify. - adjustable ultrasonic frequency (0-20hz) approx. or better,specify -possibility to adjust the impact power in percent from 10% to 100% in 10% increments - steam gas sterilizable for handpieces and probes - pedal foot switch - original trolly if available is preferred - record of data of last treatment with possibility to export data of treatments on usb key - record treatment log file. - smart integrated cooling system allowing handpiece cooling apart from suction capacity - one piece handpiece with gun shape weight -torque wrench autoclavable and reusable. |
| 6 | Suction pump | -integrated suction source with peristaltic pump. - suction activation independent from impact ultrasounds activation -gradient control of suction flow rate on the console from 10% to 100% - same foot pedal activation of suction and ultrasonic / ballistic energies -activation deactivation of energies on the console mode on /off -3 modes on the console ,hard stone ,soft stone and free mode - possibility to adjust treatment parameters ( impact /us /suction) on the console. |
| 7 | Main Power | 220 - 240 VAC , 50 Hz |